## Slide 1: Fundamentals of UV-Visible Spectroscopy
Prof. Dr. Gul Afshan Soomro
Institute of Chemistry
Shah Abdul Latif University Khairpur, Sindh, Pakistan

## Slide 2: What is Spectroscopy?
The study of molecular structure and dynamics through the absorption, emission and scattering of light.

## Slide 3: What is Light?
According to Maxwell, light is an electromagnetic field characterized by a frequency f, velocity v, and wavelength λ. Light obeys the relationship
f = v / λ.

## Slide 4
Regions of the Electromagnetic Spectrum
Light waves all travel at the same speed through a vacuum but
differ in frequency and, therefore, in wavelength.

## Slide 5
Light travelling speed:
in other media: c/n (n = refractive index, generally >1)
in a vacuum: c=2.998 x 108 m s-1 (n=1 exactly, in air n=1.0002926)
c/n= nl
Therefore:
Energy is inversely proportional to wavelength
but proportional to wavenumber
And of course, the relationship between energy and frequency:
E = hn = hc/l = hc n
h = Planck’s constant (6.626 x 10-34 J s)
n = wavenumber (most common units = cm-1)
~
~
Light is energy in the form of electromagenetic field
Review on properties of light:photon
Wavelength (l): Crest-to-crest distance between waves
Frequency (n): Number of complete oscillations that the wave makes each second
units: number of oscillations/sec or s-1 or Hertz |(Hz)

## Slide 6
Frequency Scanning Techniques: a few definitions
Emission method: source of light is sample
Absorption method: intensities of a source with and without the sample in place are compared
Spectrum: a plot of intensity vs. frequency/wavelength
In quantitative analysis:
common to work at 1 wavelength
running a spectrum is an important initial step (to select best conditions)

## Slide 7
Fig. 18-2
Regions of Electromagnetic Spectrum-the “colour” of light

## Slide 8: An Electronic Spectrum
Absorbance
Wavelength, , generally in nanometers (nm)
0.0
400
800
1.0
200
UV
Visible
maxwith certain extinction 
Make solution of concentration low enough that A≤ 1
(Ensures Linear Beer’s law behavior)
Even though a dual beam goes through a solvent blank, choose solvents that are UV transparent.
Can extract the  value if conc. (M) and b (cm) are known
UV bands are much broader than the photonic transition event. This is because vibration levels are superimposed on UV.

## Slide 9: Absorption Spectra of Hemoglobin Derivatives

## Slide 10: Components of the basic UV Visible spectrophotometer

## Slide 11: Instruments for measuring the absorption of U.V. or visible radiation are made up of the following components;
schematic diagram of a single-beam UV-Vis. spectrophotometer
Sources (UV and visible) Wavelength selector (monochromator) A Sample cell Detector A device to read out the response of the detector

## Slide 12: Sources of UV radiation
Continuous sources
Tungten lamp (350-2500 nm)
Deuterium (200-400 nm)
Xenon Arc lamps (200-1000 nm)
Line sources
Hollow cathode lamp

## Slide 13
Monochromators
Early spectrophotometers used prisms
- quartz for UV
- glass for vis and IR
These are now superseded by:
Diffraction gratings:
- made by drawing lines on a glass with a diamond stylus
ca. 20 grooves mm-1 for far IR
ca. 6000 mm-1 for UV/vis
- can use plastic replicas in less expensive instruments
Think of diffraction on a CD
http://www.mrfiber.com/images/cddiffract.jpg
http://www.ii.com/images/prism.jpg
http://www.veeco.com/library/nanotheater_detail.php?type=application&id=331&app_id=34
10mmx10mm
Why?

## Slide 14: All monochromators contain:
An entrance slit
A dispersing device (usually a prism, and gratting)
An exit slit

## Slide 15: Cuvettes
Visible; can be plastic or glass
UV; you must use quartz

## Slide 16: Transmission Characteristics of Cell Materials
Note that all materials exhibit at least approximately 10% loss in transmittance at all wavelengths

## Slide 17: Cell Types I
Open-topped rectangular standard cell (a)
and apertured cell (b) for limited sample volume

## Slide 18: Cell Types II
Micro cell (a) for very small volumes and flow-through cell (b) for automated applications

## Slide 19: Detectors
Photomultipliers are very sensitive to UV and visible radiation. They have fast response times. Intense light damages photomultipliers; they are limited to measuring low power radiation.
The photomultiplier tube is a commonly used detector in UV-Vis spectroscopy. It consists of a photoemissive cathode (a cathode which emits electrons when struck by photons of radiation), several dynodes (which emit several electrons for each electron striking them) and an anode.
A photon of radiation entering the tube strikes the cathode, causing the emission of several electrons.

## Slide 20
Photomultiplier-single channel, but very high sensitivity
- Light falls on a photosensitive alloy (Cs3Sb, K2CsSb, Na2KSb)
- Electrons from surface are accelerated towards secondary electrodes called dynodes and gain enough energy to remove further electrons (typically 4-12, to 50 with GaP).
- For 9 stages giving 4 electrons for 1, the amplification is 49 or 2.6 x 105)
- The output is fed to an amplifier which generates a signal
- To minimise noise it is necessary to operate at the lowest possible voltage

## Slide 21
Photodiode Array Spectrophotometer
- For photodiode array spectrophotometers, a white light passes through sample
- The grating polychromator disperses the light into the component wavelengths
- All wavelengths are measured simultaneously
- Resolution depends upon the distance between the diodes and amount of dispersion
No moving parts!
Simple mechanical and optical design, very compact.

## Slide 22: Conventional Spectrophotometer
Schematic of a conventional single-beam spectrophotometer

## Slide 23: Conventional Spectrophotometer
Optical system of a double-beam spectrophotometer

## Slide 24: Conventional Spectrophotometer
Optical system of a split-beam spectrophotometer

## Slide 25: Definition of Resolution
Spectral resolution is a measure of the ability of an instrument to differentiate between two adjacent wavelengths

## Slide 26: Instrumental Spectral Bandwidth
The SBW is defined as the width, at half the maximum intensity, of the band of light leaving the monochromator

## Slide 27: Natural Spectral Bandwidth
The NBW is the width of the sample absorption band at half the absorption maximum

## Slide 28: Characteristics of Beer’s Law Plots
One wavelength
Good plots have a range of absorbances from 0.010 to 1.000
Absorbances over 1.000 are not that valid and should be avoided
2 orders of magnitude

## Slide 29: Standard Practice
Prepare standards of known concentration
Measure absorbance at lmax
Plot A vs. concentration
Obtain slope
Use slope (and intercept) to determine the concentration of the analyte in the unknown

## Slide 30: Typical Beer’s Law Plot

## Slide 31: Molar Absorptivity
We have seen earlier that validation of Beer’s law is dependent on the nature of the molar absorptivity. It was found that the molar absorptivity is influenced by:
The wavelength of radiation
The refractive index and is thus indirectly related to concentration
Electrostatic interactions taking place in solution; and thus electronic distribution
In rare cases, like methylene blue, the molar absorptivity is directly dependent on concentration

## Slide 32: Molar absorptivities
    e = 8.7 x 10 19 P A
    A: cross section of molecule in cm2 (~10-15)
    P: Probability of the electronic transition (0-1)
      P>0.1-1  allowable transitions
      P<0.01  forbidden transitions

## Slide 33
Therefore, the value of the molar absorptivity is crucial for the sensitivity of an analysis.
The value of the molar absorptivity can usually range from zero (for a nonabsorbing species) to 105 (for highly absorbing species).
For quantitative analysis, a value of at least 103 is necessary for a reasonable quantitative analytical determination of an analyte
Molar absorptivities

## Slide 34
L = Cm
C = Moles/Liter
A = KCL
A = No unit C = Gram/Liter L = Cm
K=
Liter
Cm  Gram

## Slide 35: Relating Absorbance and Transmittance
Absorbance rises linearly with concentration. Absorbance is measured in units.
Transmittance decreases in a non-linear fashion.
Transmittance is measured as a %.
Absorbance = log10
  (100/% transmittance)

## Slide 37: Precision and Accuracy
Precision –
Precision +
Precision –
Precision +
Accuracy –
Accuracy –
Accuracy +
Accuracy +

## Slide 38: Spectral changes can be classed as follows:
1- bathochromic shift: The shift of λmax to longer wavelength; known by red shift, due to substitution and /or solvent effect.
2- Hypsochromic shift: The shift of λmax to shorterwavelength; known by blue shift, due to substitution and /or solvent effect.
3- Hyperchomic effect: An increase in absorption intensity (absorptivity).
4- Hypochomic effect: An decrease in absorption intensity (absorptivity).

## Slide 39: Spectral nomenclature of shifts

## Slide 40: Optically clear solvent concept

## Slide 41: Calibration methods
1- standard calibration curve
2- standard addition method

## Slide 42: UV-Vis Spectroscopy
UV- organic molecules
  Outer electron bonding transitions
  conjugation
Visible – metal/ligands in solution
  d-orbital transitions

## Slide 43: Selection Rules
In electronic spectroscopy there are three selection rules which determine whether or not transitions are formally allowed:
Spin selection rule: DS = 0
allowed transitions: singlet  singlet or triplet  triplet forbidden transitions: singlet  triplet or triplet  singlet
Changes in spin multiplicity are forbidden

## Slide 44: Selection rules
2. Laporte selection rule: there must be a
change in the parity (symmetry) of the complex.
Laporte-allowed transitions: g  u
Laporte-forbidden transitions: u  u or g  g

## Slide 45: g and u form
(g; even)
(u; odd)
d orbital are (g) and (p) orbital are u

## Slide 46
3. Selection rule of Dℓ = ± 1 (ℓ is the azimuthal or orbital quantum number, where ℓ = 0 (s orbital), 1 (p orbital), 2 (d orbital), etc.)
allowed transitions: s  p, p  d, d  f, etc.
forbidden transitions: s  s, d  d, p  f, etc.

## Slide 47: d-d transitions
Characterized by low intensities (lower molecular absorptivities ε)
ε  10-2 to102 (0.01-100)
Why?
d-d transition are Laporte-forbidden transitions
Involve g-g transition (d orbital are g)

## Slide 48: Ligand to metal charge transfer complex (LMCT)
a. intramolecular e- transfer from ligand-based orbitals to metal-based orbitals.
Mn+___L- M (n-1)+____L
b. ε values are 5000-50,000 M-1 cm-1
(very intense)
Since we are “oxidizing” the ligand and “reducing” the metal, LMCT are favored with
i.Easily oxidized ligands (-donors)
ii. Easily reduced metals (i.e. High oxidation state)

## Slide 49: Ligand-to-Metal Charge Transfer (LMCT)
Occurs with metals in high oxidation states (few d electrons) complexed by ligands which are good electron-donors
Permanganate anion
Manganese(VII) has no d electrons, while the O2- (oxide) ligands donate lone pairs of electrons to form double bonds

## Slide 50: UV-Visible spectrum of MnO4-
Lp  M(t2g)
Lp  M(eg)

## Slide 51: Metal to ligand charge transfer complex (MLCT)
a. intramolecular e- transfer from metal-based orbitals to ligand-based orbitals.
M___L M +____L –
b. Similar to LMCT in terms of intensities.
c. Not as common as LMCT.

## Slide 52: Metal-to-Ligand Charge Transfer (MLCT)
Ru(2,2’-bipyridine)32+
Occurs with metals in low oxidation states (lots of d electrons) complexed by ligands which are good electron-acceptors
Ruthenium(II) has 6 d e-
2,2’-bipyridine is a good p-acceptor ligand

## Slide 53: Visible spectrum of Ru(2,2’-bipyridine)32+
[Ru(bpy)3]2+ MLCT
lmax = 452 nm e = 14,600 M-1 cm-1

## Slide 54: Absorption by Inorganic Groups
Inorganic groups containing double bonds absorb in the UV-Vis region. The most transitions are a result of n-p* transitions as in nitrate (313 nm), carbonate (217 nm), nitrite (280 and 360 nm) and azide (230 nm)

## Slide 55: Absorption Involving d and f Orbitals
Many transition metals have colored solutions and are also colored in the solid state.
The transition metals have some of their d orbitals empty where a d-d transition can occur.
The d-d transitions require excitation energy in the UV-Vis region. The direct interaction of the d electrons with ligands around the transition metal results in a spectrum of broad band nature.
On the other hand, inner transition elements show transitions by absorption of UV-Vis radiation (f-f transitions).

## Slide 56
Since the electrons in the f orbitals are far inside the metal orbitals and are screened by electrons in orbitals of higher principal quantum numbers, f-f electronic transitions will not be affected by the nature of ligands or solvent around the inner transition metals.
Therefore, the spectra of inner transition metals have narrow bands.
Absorption Involving d and f Orbitals

## Slide 58
The crystal field theory is usually used to explain splitting of the d orbital energy so that a transition from a lower energy d orbital electron can be excited to a higher energy d orbital.

## Slide 59: Charge Transfer Absorption
When a ligand permanently donates an electron to a metal ion, a charge transfer is said to take place.
The net outcome of the process is an oxidation reduction phenomenon occurring within the complex.
An example is the reaction of Fe3+ with thiocyanate where the product is an excited species with neutral thiocyanate and Fe3+.

## Slide 60
In less common situations, the transfer of electrons can take place from the electron deficient metal ion to the ligand.
An example is the Fe2+ or Cu+ complexes with 1,10-phenanthroline where Fe2+ and Cu+ metal ions donate electrons to 1,10-phenanthroline. The complex will then have Fe3+ and Cu2+ ions.
Charge transfer complexes are of special interest their molar absorptivities are unusually high; allowing very sensitive determinations.

## Slide 62: Absorption in the UV-VIS Region
Chemical species that can form soluble complexes with chromophoric reagents.
Examples
   thiocyanate ion for iron, cobalt, & molybdenum;
   diethyldithiocarbamate for copper;
   diphenylthiocarbazone for lead;
   1,10-phenanthroline for iron.

## Slide 63: Limitation of UV-visible spectroscopy
1. Deviation from beer´s law
Positive deviation
Negative deviation
2. Instrumental deviation
Noise originating from the light source and electronicComponents:
Baseline flatness as the intensity of the radiation coming
from the light sources is not constant over the whole UV-Visible range.
3. Chemical deviation
mixtures of molecules can be a problem due to overlap (hence, routinely requires significant sample preparation)

## Slide 64
Thanks